UVIT detector system Engineering Model test and calibration.

J. Postma and J. Hutchings

University of Calgary, and NRC Herzberg Institute

This report describes the test and calibrations carried out at Routes, DFL, and the University of Calgary, in October - December 2008. The scope and aims of the work are laid out in the calibration test plan. The data are stored and available for further work, and the test details are recorded in spreadsheets. The results and images in this report were generated during the calibration period and subsequently.  The EM was shipped to IIA and tested there in January 2009.  The EM uses an NUV photocathode, and all work was done in the wavelength range 190-270nm.

The box below gives the overall summary of the work.  This draft March 23, 2009.

Figures 1 to 4 show the Calgary vacuum facility used, and photos of the hardware in the facility.  Figure 5 shows the GSE display, in the calibration facility area.

Fig 1. Schematic of the Calgary facility. Only the UV arm was used, with a deuterium lamp.

Fig 2. REA box in clean area outside tank

Fig 3. Vacuum tank with lid slightly raised. UV light channel is on the right.

Fig 4. EM detector and HVU (below) inside the tank, on the movable mounting platform.  Light port is on the right.

Fig 5. GSE control and display screen.

DFL data

We first show data obtained at DFL during the TVAC testing. Figure 6 shows dark frames (entire CMOS image raw readouts), displayed as plots of all 512 individual pixel rows.  The initial row is offset to a lower value, but the rest lie at approximately the same level.  There is a signal ramp along the rows, rising to the right, which is present in all data. The photon events are superposed on this, and proper event centroiding requires subtraction of the local background. This is measured, as described in Postma et al 2007, as the minimum corner value in the 5 x 5 pixel box centred on each photon event.

Figure 6 displays the dark pixel rows at cold and hot temperatures, and gives some statistics of the images.  The behaviour does not change significantly with temperature, but a number of `hot’ pixels are seen in the hot images.  The hot pixel values are much lower than the threshold values used for event recognition, and do not affect the operation in any way.  The dark images are referred to as background or bias images.  Figure 6 shows the darks for 3 different gain settings.  These scale well with gain and show no problems for use in photon counting.

Fig 6. Dark frames at cold and hot temperatures.

Fig 7. Dark frame data at 3 different gain settings.  Normal operations will be at settings 1 or 2. 8 is the maximum.

Fig 8. Dark single frame on left. Single frame with uneven light illumination on right.

Figure 8 shows individual dark and light frames. The dark frame contains single bright pixels that will not be identified as photon events, which are seen in the light frame.

Fig 9. Left: event total signal for 3 algorithms. Right: Event maximum pixel signal.

There are 3 templates for event centroiding, as described by Postma et al (2007).  They are designated 3cross (3C), 3square (3S), and 5square (5S). Figure 9 shows the total signal for the same events, within the three templates.  As most of the event signal lies within the 3C template, little more is added in the others. There is a spread of event signals, and a threshold value is set by inspection of these distributions, which also contain a spike of low energy `events’ which are noise and not real photons.

Events are recognized as maximum pixel values surrounded by 8 lower values, and a separate threshold is set for this value, from the distributions in the right panel of Fig 9, again in order to eliminate low signal non-photon events.  These events are shown with gain setting 1.  Gain setting 2 will double the signals. As the A/D converter saturates at 128000, gain 2 is useable here. Should the detector response drop during use or ageing, use of higher gain may be helpful.

Fig 10. Flat fields with the three different templates – L to R: 3C, 3S, 5S

Figures 10,10b shows the resulting image accumulating many frames of data with uniform illumination of the detector. The nature of the centroiding algorithms leads to systematic errors, which underpopulate the CMOS pixel boundaries, as discussed by Postma et al, and demonstrated well in these EM data.

While corrections are possible for the 3C and 3S images, the 5S results give a fairly good result without correction. The value of the smaller templates lies in being able to detect more photons separately in crowded or bright fields of view, and will be selected beforehand, with the scientific aims in mind.  All algorithms use the corner minimum of 5S for background level estimation.

Illumination with low and high count-rate sources was done to measure these. The low count rate source was fibre optic UV feed into chamber, not shining directly on to detector. The high count rate source was DFL’s vacuum chamber mass spectrometer, likely shining directly on to detector.

| Algorithm | Low Rate1 | High Rate2 |
| --- | --- | --- |
| 3x3 Cross | 19.72 c/f | 1411.58 c/f |
| 3x3 Square | 21.82 c/f | 1356.91 c/f |
| 5x5 Square | 21.63 c/f | 1215.40 c/f |

- At this low of count rate, the lower rate for the 3x3 Cross is likely due to the shape energy being less for that algorithm, thus less able to satisfy the energy threshold.  In the plots of Event Count vs. Frame Number below, you can see the cosmic hits superimposed.  Cosmic counts rate in space can be calibrated in orbit, and as these should be relatively constant, can be corrected out of the photometry.   Also, all voltages were run at nominal settings:  Anode = 5000V, MCP = 2020V, Cathode = -400V, full frame (512x512) at 28.7Hz.

- The MCP voltage was lowered to 1950V for this bright source (mass-spec).  Imaging at 2020V would BOD soon after the voltage ramped.  Cosmic hits are not noticeable at this count rate, but in space the measured cosmic count rates from in-orbit calibration viewing a dark (background) sky can be used for photometric correction. We describe low MCP voltage operations in more detail later.

Fig 10b. 1/32 pixel resolution cuts through 3C-centroid flat fields.  X-cut on the left and Y-cut on the right.

Fig 11. 3C images at low count rate, clockwise: image at 1 pixel resolution; minimum corner values (i.e. the bias image); number of photon events per frame for 55000 frames; the max-min corner values (high points indicate double-photon events).

Low MCP voltage

Reducing the MCP voltage reduces the signal gain, and this allows exposure to brighter light without draining the MCP charge and shortening its lifetime.  We did extensive calibrations of this, reported in the light-spot data later in the report. Figure 12 shows the reduction on whole-field counts under uniform low level illumination, with MCP voltage variations. This offers an order of magnitude extension of dynamic range.

Fig 12. Photon events counted as function of MCP voltage, under low level flat-field illumination. As there is little crowding of events, all three algorithms have the same response. Nominal operating voltage is 2020.

Figure 13 shows the spreads of event signals for the range of MCP voltages used from the full-field low illumination case. These plots show only the events above the threshold that eliminates non-photon faint events.

Fig 13.  5S Event total signal for range of MCP voltages.

Calgary lab tests and results

Fig 14. Statistics of run of raw frames. This one is for dark.

In the Calgary facility, it was not possible to illuminate the whole field (an integrating sphere has since been obtained for the FM calibrations). Some initial runs were done with dark and some illumination to accumulate and examine statistics of the type shown in Fig 14. These were run at gain values 1, 2, 4, and 8. No anomalies were seen and the data reproduced those from the tests at DFL.

The bulk of the work was done with illumination through a hole mask (see fig 15), with holes alternating through 25, 50, and 75 micron diameters. The mask was placed at 3 different distances from the detector, and diffraction of the incident collimated beam produced 3 different sizes of image for each hole. The hole sizes were not resolved, but served to cover a range in signal level.  In the largest distance, the spots were many pixels across, and resemble small galaxies as will be observed in flight. At the smallest distance (the mask was held against the photocathode window), the spots are comparable with the pixels, and can be used to infer the spatial resolution of the detector.

The monochromator was run at 230nm for all tests, and also at 190 and 270 for some, to test wavelength sensitivity. The light level was adjusted via an aperture to the lamp, and measured by a photometer viewing the light from a beamsplitter in the beam.  Light levels ranged from counting all photons, to significant saturation of the photon-counting, to measure dynamic range.

All setups were run using all 3 centroid algorithms, and integrate mode.  We also ran subarray readouts at higher rates, to calibrate that for dynamic range.

Fig 15. Image through hole mask at intermediate distance. The 3 holes sizes give rise to the different brightness levels.

Fig 16. Intermediate size hole images seen with different centroid algorithms, anticlockwise from top left: 3C, 5S, 3C summed and 5S summed.  The summed images combine about 6 spots with different readout pixel placement, so the pixel edges are smoothed out. The spots are aligned in the vertical direction, so this smoothing is noticeably less effective in X. In flight, spacecraft drift will give this smoothing, as the spots are shifted to compensate for the drift in final science image construction.

Small spots

The smallest spots have FWHM comparable with a readout pixel, and are close to the expected point source images through the flight telescope.  They may be used to estimate the detector resolution, and also to calibrate saturation effects and double-photon rejection thresholds.

Fig 16. Small spots in photon-counting and integrate mode.

Integrate mode simply adds all the photon events without centroiding, and so produces spots that have the width of the event splashes on the CMOS. They may be centroided from the image, to sub-pixel precision, and used for monitoring the spacecraft drift by following one star in the visible channel.

Figure 16 illustrates this difference for the smallest spots.

The table below outlines the grid of settings used for imaging with the smallest spots (mask against the photocathode window).

Fig 17. Small spot sizes  in mask and on detector.

In Figure 17 we show the measured FWHM of our smallest spots. The relationship with the mask hole diameters is close to linear, so if we extrapolate this line to zero size spot, we have a measure of the detector system spatial resolution. This is 1.3 pixels, which will correspond to 1.1” on the sky in UVIT.  It is of interest to note that the resolution measured at Photek for this detector module (gap spacing and window voltage), corresponds to 1.13 arcsec on the sky, so we appear to have realized the full resolution expected for this detector. The flight units have smaller window gaps and thus should give better resolution than the EM.

Fig 18. Saturation effects

We measured the saturation of the photon-counting by increasing the light level and measuring the detector signal. Figure 18 shows this for two spots of different light level (large and small in our hole mask), and the saturation can be seen clearly, more in the brighter spot, as expected.  These results use all reported event centroids, and clearly increasingly include double or multiple photon events with a single read frame.  While the photometric correction can be estimated by this kind of plot, double events will give poor centroids and degrade the spatial resolution. We discuss below how we can reject double events to avoid this.  We also note that the results here are the same for all centroid algorithms, because the spots are well separated in a dark field. In flight, observations with a crowded or bright background, the 3C and 3S algorithms will see more of the photons from sources that are close to each other, than the 5S.

We can mitigate the saturation in two ways – fast readout of subarrays, or lower MCP gain. With the smallest subarray we can read 20 times faster, and thus reduce the photon events per frame by this factor. The effect in Fig 18 is to stretch the saturation curve by this factor to brighter fluxes, and the plot shows this – the brightest objects are still showing linear response to light levels.

Turning to lowering the gain via MCP voltage, we get the second linear relationship in Fig 18.  Here we have reduced the photon events by lowering the gain so the events fall below the threshold, into the noise. This is less than the factor 10 from figure 12 for flat field illumination, so the results depend on how crowded or bright the background is of the object(s) of interest.

Low MCP voltage effects on photon events

Fig 19. Effects of MCP voltage on spot image.

Figs19, 20, 21, 22  show another important effect of MCP voltage.  The image becomes smaller as well as fainter, and suggests that the spatial resolution can be improved by this low gain mode. The spot sizes scale linearly as they do for the full voltage images in Fig 17, and the extrapolation to zero spot size suggests a resolution of 0.6”, down from 1.1”. We are still investigating whether this is real, and have designed tests for the FM units that will clearly show this.  Since we are detecting only the most energetic photon events, it may be that the electron spread on to the MCP is smaller. However, if we set the threshold very high at nominal voltage, we do not get the same resolution improvement.

Fig 20. Stacked raw frames of the same images with different MCP voltage.

Fig 21. Photon event changes with MCP voltage.

Fig 22.  Photon events at different MCP voltages

Fig 23. Flat fields at MCP voltage 2020 (left) and 1800 (right)

Figure 23 shows flat fields (5S photon-counting images) at nominal and low MCP voltage. The pixel-to-pixel noise is higher in the low voltage case, and low frequency structure is also more apparent. Once again, we are seeing the most energetic events only at low MCP voltage.  If we are to use lower MCP voltages for science, we will need to calibrate a few settings well.

Double-photon rejection

We may reject double photon events in constructing the image by using the max-min corner difference value from the 5x5 event subimage. Two photons not exactly aligned will produce a wider double-event for centroiding, and also spill signal into one of the corners.  Figure 24 shows how this affects one small spot with some structure. As we raise the rejection threshold, we lose signal, but the signal that remains is composed of single events that give better image resolution.

Fig 24. Double-photon rejection in small image (1/8 pixel bins).

Figure 25 shows the largest spots, with the mask at greatest distance from the window. The diffraction effects of the 3 different apertures can be seen clearly. These are of interest as they resemble extended galaxy images with structure over many readout pixels, and we can examine the effects of high flux and double-photon rejection.

Fig 25. Large spots with diffraction, from hole mask at large distance from the detector.

Fig 26. Effects of double-photon rejection on large spots of different signal level.

In Figure 26, the rejection thresholds applied are the numbers shown (max-min corner values for each event), and the images are constructed without events which have values above the thresholds. This type of image construction can be done on the ground with as many different thresholds as wanted, by the scientist, so no choice is required before making the observations. It does need to be an option in the data products and analysis tools provided.  Figure 27 shows the detailed effects of the threshold on a bright source with many double events.

Fig 27. Effects of rejection threshold on large bright source.

Figure 28 summarizes the effects of the threshold on the total signal for bright sources, in the left panel.  As expected, the effect is greatest in bright sources where there are many double events.  The right panel shows the ratio of signal using the 5S to the 3S algorithm for centroiding. In bright objects with many double events, the smaller algorithm is more efficient.

Fig 28. Statistics of rejection threshold choices.

References

Hutchings J.B., Postma J., Asquin D., Leahy D., PASP, 119, 1152, 2007

Postma and Hutchings EM calibration plan

Routes, CSA, et al  FM contamination plan

Summary and way forward

The bullets below summarize the EM calibration results and outline the basis for FM calibration planning.  The EM has been delivered to IIA in Hoskote and operated successfully there. Since then, a cleanliness plan for the FM units has been developed and is in place. We have also had discussions on FM contamination control procedures after delivery to IIA and ISRO.  We thank Routes and CSA/DFL for our participation in these tests.

- EM works very well, and is up to flight specs

- No significant dark current issues

- Resolution as expected from Photek tests

- Rapid read small window works well

- Low MCP voltage usable mode

- Bright limit extended 6 mags with these procedures

- BOD protects system during MCP ramp-up, as designed

- Integrate and all centroid modes work well

- Photon arrival saturation calibratable over factor 4

- Double-event rejection works well

- Flight spatial resolution will depend on telescope optics

Planning goals for FM calibrations

- Integrating sphere for flat fields, geometrical distortion,  now in place

- Calibrated diode for absolute flux calibration

- Flight filters and gratings to be included

- Characterize resolution, event size and energy stats

- Calibrate at selected low MCP voltages

- Set operating thresholds and voltages

- Set operating row gap and gain

- New hole mask with pairs of apertures for resolution measurement

- Further hole mask with large holes for edge focus tests

- Server acquired for data storage and distribution

- Air shower and flow bench to be installed

- Witness sample transmission testing in place

# Text boxes

2020  1960

1900 1850

1800

MCP volts